# Teleophthalmology Review for AFIO

# 1. Device Categories Overview

Portable fundus imaging devices fall into four categories based on cost, image quality, and integration complexity:

| Category | Examples | Cost Range | Img Quality | Integration |
| --- | --- | --- | --- | --- |
| DIY / 3D Printed | 25D lens adapter, GSoP, DIYretCAM | $10 – $200 | Moderate | Manual |
| Low-Cost Commercials | Jaiz RetiCAM, D-EYE, Peek Retina, Volk iNview | $100 – $600 | Good | WiFi/USB |
| Chinese Medical Grade | Hongdee FC-161, FC-3100, Lety Medical | $800 – $6,000 | Very Good | WiFi/DICOM |
| Premium Validated | Remidio FOP NM-10 | $7,400+ | Excellent | WiFi/DICOM |

# 2. DIY and Ultra Low Cost Devices

## - 25D Condensing Lens Adapter

What it is: A 25-diopter Volk or equivalent condensing lens combined with a smartphone camera and a 3D-printed or cardboard housing. The smartphone LED flash provides illumination. No electronic components other than the smartphone itself

Image quality: High-resolution 50° field of view. Published studies report images comparable to standard non-mydriatic fundus cameras for central retinal pathology. Requires dilated pupils (mydriasis) for reliable peripheral imaging

Cost: $30 – $100 for the lens. 3D-printed housing adds $5–$20. Total under $120 (~34,000 pkr)

Transfer method: Smartphone camera roll → manual share via WhatsApp, email, or upload to web app. No automatic transfer. Manual step required by technician

Diseases detectable: DR (Grade 2+), HR, Glaucoma (optic disc assessment), AMD, retinal detachment, anything visible in the central 50° of the retina

Integration with AFIO ScanAI: Technician opens AFIO ScanAI web app on smartphone browser, uploads the captured image directly. System runs inference and generates report. Simple and functional for a pilot.

Limitations: Requires dilated pupils. Image consistency depends on technician skill. No standardized illumination. Not suitable for high-volume clinical use.

Verdict: Best for proof-of-concept and research validation. Not recommended for sustained clinical deployment

## - GSoP, Glaucoma Screening on Phone (Research Prototype)

What it is: A 3D-printed smartphone fundus imaging adapter developed using Solid Edge 2022 design software and a Prusa I3 MK3S 3D printer. Uses a beam splitter, reflective mirror, and transparent paper as optical elements. Designed specifically for optic disc imaging for glaucoma screening.

Cost: Under €10 for components. 3D printing adds ~$10. Total cost well under $25 (~ 7,000 pkr)

Image quality: 32° field of view. Captures non-mydriatic color and 3D retinal images. Published study showed glare artifacts from the beam splitter that reduced diagnostic image quality compared to commercial devices

Transfer method: Smartphone video/photo → manual upload.

Diseases detectable: Glaucoma (optic disc region specifically). Limited for DR or peripheral pathology due to 32° FOV.

Verdict: Interesting research device but not clinically validated enough for deployment

## - DIYretCAM (Raspberry Pi Based)

What it is: A portable non-mydriatic fundus camera built using a Raspberry Pi computer, infrared-sensitive camera board, dual infrared and white light LED, 20-diopter condensing lens, and a 5-inch LCD touchscreen. Fits in a white coat pocket at 133mm × 91mm × 45mm.

Cost: $185.20 to build including the lens. All components available from standard electronics suppliers (~ 50,000 pkr)

Image quality: 5–6 disc diameter field of view. Non-mydriatic. Infrared pre-illumination reduces pupil constriction. Published image quality sufficient for optic disc assessment.

Transfer method: USB, WiFi via Raspberry Pi. Images saved as JPEG. Can be configured to auto-upload to a server.

Integration with AFIO ScanAI: Raspberry Pi can be programmed to automatically send captured images to your AFIO ScanAI server via HTTP POST. No manual step needed if configured correctly.

Diseases detectable: Glaucoma (optic disc), central retinal assessment. Limited field of view restricts peripheral DR assessment.

Verdict: Best DIY option for a technically capable team. MCS students could build this. Good for glaucoma-focused screening but not recommended for full DR screening due to limited FOV

# 3. Low-Cost Commercial Devices

## - Peek Retina

Made in: UK (Peek Vision, London). Designed specifically for low-resource settings.

What it is: A smartphone fundus adapter designed for mass population screening in sub-Saharan Africa and South Asia. Very simple to use, non-specialist staff can operate after minimal training

Cost: ~$500 USD (Rs. 140,000 approximately)

Image quality: 40° field of view. Sensitivity for DR: 84%, Specificity: 79.9% compared to standard fundus camera. Good for glaucoma optic disc assessment (kappa 0.69 agreement with standard camera).

Transfer method: Peek companion app → cloud sync → accessible via web dashboard. Has its own telemedicine platform.

Diseases detectable: DR, Glaucoma, HR, clinically validated for all three.

Integration with AFIO ScanAI: Images can be exported from the Peek app as JPEG and uploaded to AFIO ScanAI

Verdict: Good device for population-level screening. The built-in Peek telemedicine platform means it comes with its own ecosystem.

## - Jaiz RetiCAM 2.0

Made in: India (Jaiz Aris Medical Technologies, Bangalore)

What it is: A smartphone-based fundus imaging adapter compatible with all smartphones. Uses a 20D lens. Can be mounted on a slit lamp for stability or used handheld for outreach camps. Has been in clinical use since 2016 with ophthalmologists globally.

Cost: Rs. 20,000 (~$70 USD) per unit. Cheapest validated clinical option available

Image quality: Comparable to direct ophthalmoscopy. 40° field of view when used correctly. Requires dilated pupils for best results. Has an app that automatically selects the best frame from continuous video capture.

Transfer method: Smartphone camera → Jaiz companion app → WiFi/USB share. Manual upload to AFIO ScanAI.

Diseases detectable: DR, HR, Glaucoma (disc), AMD, all conditions visible on central retinal fundus photography.

Integration with AFIO ScanAI: Technician opens AFIO ScanAI on smartphone browser, selects image from Jaiz app gallery, uploads. Results shown in 3 seconds

Shipping to Pakistan: Ships from India, 3–5 days via courier (DHL/FedEx). No customs complications as medical import from India is straightforward. Total landed cost approximately Rs. 25,000–30,000 per unit.

Verdict: BEST VALUE OPTION. Rs. 20,000 per unit, 40° FOV, clinically used since 2016, ships from India in days

## -  D-EYE

Made in: Italy (D-EYE SRL, Padova). FDA cleared.

What it is: A small magnetic clip-on adapter for iPhone that converts it into a digital direct ophthalmoscope. Uses cross-polarization optical design to reduce corneal reflections. Originally designed for iPhone 5/6 but newer versions support current models.

Cost: Approximately $390 USD (Rs. 110,000 approximately)

Image quality: 20° field of view, smaller than other devices. Sensitivity for DR diagnosis varied between 55–96% in published studies. Lower consistency than Remidio or Jaiz for DR specifically.

Transfer method: D-EYE app on iPhone → share via standard iOS share sheet. WiFi/USB.

Diseases detectable: Glaucoma (disc assessment), central DR, HR. Limited peripheral reach due to 20° FOV.

Shipping to Pakistan: Ships internationally from Italy. 7–14 days. Import duty applies. Landed cost approximately $450–500.

Verdict: Higher cost than Jaiz with lower FOV and lower DR sensitivity. iPhone-only is a limitation

## - Volk iNview

Made in: USA (Volk Optical, Mentor Ohio)

What it is: A wide-angle ophthalmic camera for iPhone that acquires 50° digital color images of the retina. App automatically detects and selects the most focused images from video capture. Non-mydriatic.

Cost: ~$500 USD. (Rs. 140,000 approximately), iPhone not included.

Image quality: 50° field of view, best among low-cost devices. Non-mydriatic. Auto-image selection from continuous video.

Transfer method: Volk iNview app → iPhone camera roll → share. WiFi/USB.

Diseases detectable: DR, HR, Glaucoma, AMD, wide FOV allows peripheral pathology detection.

Verdict: Best image quality in this price range. iPhone-only is a significant limitation

# 4. Chinese Medical Grade Devices

China is the global manufacturing hub for medical-grade ophthalmic equipment. Devices from manufacturers like Beijing Hongdee Corporation offer clinical-grade image quality at 30–70% lower cost than equivalent Western devices. All devices carry CE certification and many have FDA clearance as well

## Beijing Hongdee Corporation, Overview

Hongdee has manufactured ophthalmic equipment for 15+ years. Product range includes fundus cameras, OCT, slit lamps, tonometers, and auto perimeters.

Certifications: CE, FDA, ISO 9001, ISO 13485, SGS.

Shipped to 60+ countries. Ships within 2–7 days after order with 12-month warranty and 10-day unconditional return policy.

Website: https://www.hongdeecorp.com/

## - Hongdee FC-161, Recommended Portable Option

What it is: Portable handheld non-mydriatic fundus camera. 12MP camera with 4.3-inch touchscreen. Manual and auto focus. Takes photos and videos. Built-in rechargeable battery.

Transfer method: WiFi and USB, both supported. Connect to hospital WiFi and transfer directly to server, or connect USB cable to laptop and copy files.

Image quality: 12MP, 40–45° field of view. Non-mydriatic. Comparable to clinical-grade portable cameras from Optomed at a fraction of the cost.

Cost: $1,500 – $2,500 USD (Rs. 420,000 – 700,000).

Diseases detectable: DR, HR, Glaucoma, AMD, all retinal conditions visible on standard fundus photography.

Integration with AFIO ScanAI: Image captured → transferred via WiFi to shared folder or directly uploaded to AFIO ScanAI web interface. System runs DR/Glaucoma/HR inference and generates report automatically.

Shipping to Pakistan: Air freight from Beijing to Islamabad via DHL or FedEx. 3–5 transit days. Shipping cost ~$150–300 per unit. Customs duty 5–20% on medical equipment HS code 9018.

Estimated landed cost: Rs. 500,000 – 850,000 per unit all-in including shipping and duty.

Verdict: Best mid-range option. Professional image quality, WiFi transfer, non-mydriatic. Recommended

## - Hongdee FC-3100, Best Integration Option

What it is: Desktop non-mydriatic fundus camera. 24MP sensor. Fully automated, motorized chinrest adjusts patient head position automatically. Small pupil mode works with pupils as small as 3.3mm. 9,216 internal fixation targets.

Transfer method: DICOM 3.0, USB, LAN, and FTP. The DICOM 3.0 support is the key feature, images can be sent directly to your AFIO ScanAI DICOM receiver the moment a scan is taken, fully automatic.

Image quality: 24MP, auto focus and capture. Highest image quality in Hongdee lineup. Comparable to Topcon non-mydriatic cameras.

Cost: $4,000 – $6,000 USD (Rs. 1.1 – 1.7 million). For 30+ units price drops to $3,236 per unit.

Diseases detectable: DR, HR, Glaucoma, AMD, Retinal detachment, full posterior segment assessment.

Integration with AFIO ScanAI: FC-3100 configured with AFIO ScanAI server IP and DICOM port. Every scan is automatically sent to your system via DICOM. No manual step. Your DICOM receiver processes it, runs inference, and generates the report, fully automated end to end.

Shipping to Pakistan: Same as FC-161. Air freight $150–300. Customs duty 5–20%.

Estimated landed cost: Rs. 1.2 – 1.9 million per unit all-in.

Verdict: Best technical integration. DICOM 3.0 makes the workflow fully automatic. Recommended for AFIO

# 5. Premium Validated Device

## - Remidio FOP NM-10

Made in: India (Remidio Innovative Solutions, Bangalore). FDA cleared December 2025. CE marked.

What it is: The world's first smartphone-enabled non-mydriatic cloud-synced retinal imaging system. Patented optics. Guided auto-capture operable by non-specialist staff with no ophthalmology training. Used in mass screening programs across South Asia and Africa.

Image quality: 45° field of view. 8 internal fixation points. Non-mydriatic. Sensitivity for DR: 93%, Specificity: 98%. Referral DR sensitivity: 88%, Specificity: 95%. Best image quality of any smartphone-based device.

Transfer method: WiFi, images automatically sync to Remidio cloud. Can also export JPEG and PDF. DICOM export available.

Cost: €6,900 (regular price €9,900), approximately Rs. 2.1 million.

Diseases detectable: DR, Glaucoma, HR, AMD, all three conditions your models screen for, plus AMD. Validated in published studies for all conditions.

Integration with AFIO ScanAI: Images exported as JPEG from Remidio app uploaded to AFIO ScanAI. Or DICOM export sent directly to your server. Remidio also has its own cloud platform, you would need to decide whether to use theirs or yours.

Shipping to Pakistan: Ships from India. 5–10 days. Import relatively straightforward. Landed cost approximately Rs. 2.2–2.4 million per unit.

Verdict: Best clinical validation of any portable device. FDA cleared. At Rs. 2.1 million per unit it is expensive for mass rollout. Recommended only for highest-volume rural sites or demonstration purposes.

# 6. Full Device Comparison

| Device | Origin | Cost (USD) | FOV | Non-Mydriatic | Transfer | DICOM | DR Valid. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 25D Lens DIY | DIY | < $120 | 50° | No | Manual upload | No | Moderate |
| GSoP | DIY | < $25 | 32° | Yes | Manual upload | No | Limited |
| DIYretCAM | DIY | $185 | Limited | Yes | WiFi/USB | No | Limited |
| Jaiz RetiCAM | India | ~$70 | 40° | Partial | WiFi/USB | No | Clinical |
| D-EYE | Italy | $390 | 20° | Partial | WiFi | No | 55–96% |
| Peek Retina | UK | $500 | 40° | Yes | Cloud | No | 84% |
| Volk iNview | USA | $500 | 50° | Yes | WiFi | No | Clinical |
| Hongdee FC-161 | China | $1,500–2,500 | 45° | Yes | WiFi+USB | No | Clinical |
| Hongdee FC-3100 | China | $4,000–6,000 | 45° | Yes | DICOM+LAN | Yes | Clinical |
| Remidio FOP | India | €6,900 | 45° | Yes | WiFi+DICOM | Yes | 93% |

# 7. Integration With AFIO ScanAI

## - Manual Upload Pathway (Jaiz, D-EYE, Peek, Volk, 25D Lens)

For devices without DICOM support, the integration pathway is manual but simple:

- Technician captures fundus image on device/smartphone

- Opens AFIO ScanAI web app in smartphone browser

- Uploads image via the New Screening workflow

- Selects conditions to screen (DR, Glaucoma, HR)

- System runs inference in under 3 seconds

- Report generated and saved to MongoDB database

- Accessible to all roles through the dashboard

No special hardware or network configuration needed. Works over any internet connection including 4G mobile data

## - WiFi Transfer Pathway (Hongdee FC-161)

The FC-161 connects to hospital WiFi and can transfer images to a shared network folder automatically. Your system watches this folder and processes new images:

- FC-161 configured with hospital WiFi credentials

- After each capture, image automatically saved to shared network folder

- AFIO ScanAI folder-watch service detects new file

- System reads patient ID from filename or prompts technician

- Inference runs automatically

- Report generated and technician notified via dashboard

## - DICOM Pathway (Hongdee FC-3100, Remidio FOP)

DICOM integration is fully automatic with zero manual steps after initial configuration:

- FC-3100 or Remidio configured with AFIO ScanAI server IP and DICOM port (typically 4242)

- Every scan automatically sent via DICOM C-STORE to your server

- Your DICOM SCP listener receives the file

- Patient demographics extracted from DICOM tags automatically

- Image extracted and inference runs automatically

- Report generated, saved to MongoDB, all roles notified

- No technician action required after the scan

This is the most clinically appropriate integration

# 8. Importing Devices to Pakistan, Procedure

## - Import Procedure for China (Hongdee)

- Contact Hongdee directly via hongdeecorp.com or WhatsApp

- Request proforma invoice and CE/FDA certificate for customs

- Payment via T/T bank transfer or Alibaba Trade Assurance

- Hongdee ships via DHL/FedEx air freight from Beijing

- Package arrives Islamabad International Airport in 3–5 days

- Customs clearance: HS code 9018 (medical instruments)

- Import duty: 5–20% depending on exact classification

- Clearing agent in Islamabad handles customs: Rs. 5,000–15,000 fee

- Total process from order to delivery: 11–22 days

## - Import Procedure for India (Jaiz RetiCAM, Remidio)

- Order directly from Jaiz (jaizaris.com) or Remidio (remidio.com)

- Payment via wire transfer or credit card

- Ships via standard courier, DHL or BlueDart India to Pakistan

- Transit time: 3–7 days

- Customs generally simpler for India-origin medical equipment

- Total process: 7–14 days

## Estimated Landed Costs in Pakistan

| Device | Device Cost | Shipping | Duty (est.) | Clearing | Total (PKR) |
| --- | --- | --- | --- | --- | --- |
| Jaiz RetiCAM 2.0 | Rs. 20,000 | Rs. 5,000 | Rs. 3,000 | Rs. 5,000 | ~Rs. 33,000 |
| D-EYE | Rs. 109k | Rs. 15,000 | Rs. 20,000 | Rs. 8,000 | ~Rs. 152,000 |
| Peek Retina | Rs. 140k | Rs. 15,000 | Rs. 25,000 | Rs. 8,000 | ~Rs. 188,000 |
| Volk iNview | Rs. 140k | Rs. 20,000 | Rs. 25,000 | Rs. 8,000 | ~Rs. 193,000 |
| Hongdee FC-161 | Rs. 560k | Rs. 50,000 | Rs. 90,000 | Rs. 12,000 | ~Rs. 712,000 |
| Hongdee FC-3100 | Rs. 1.4M | Rs. 80,000 | Rs. 250,000 | Rs. 15,000 | ~Rs. 1.74M |
| Remidio FOP NM-10 | Rs. 2.1M | Rs. 50,000 | Rs. 350,000 | Rs. 15,000 | ~Rs. 2.5M |

# 9. Recommendations

## - Phase 1, Pilot (immediate, low cost)

Use Jaiz RetiCAM 2.0 at Rs. 20,000 per unit. Order 3–5 units, ships within 7 days. Equip technicians at 3–5 rural CMHs. Images uploaded manually to AFIO ScanAI via smartphone browser. This validates the workflow end-to-end at minimal cost before investing in more expensive hardware.

- Cost: Rs. 33,000 per site (all-in)

- Setup time: 1 week including technician training

- Risk: Low, cheap enough to replace if damaged

## - Phase 2, Scale (3–6 months after pilot)

If pilot succeeds, upgrade to Hongdee FC-161 for higher volume sites. Professional image quality, WiFi transfer, 12MP non-mydriatic. Better suited for 20+ patients per day.

- Cost: ~Rs. 712,000 per site

- Appropriate for district hospitals and larger CMHs

## - Phase 3, Full Integration (12 months)

Deploy Hongdee FC-3100 with DICOM 3.0 at AFIO main site and major partner hospitals. Fully automatic workflow, scan taken, image appears in AFIO ScanAI, report generated, all roles notified. No manual step.

- Cost: ~Rs. 1.74M per site

- Full DICOM integration with your system

- Supports future PACS integration